## Slide 1
HeartRateAnalyze
ESP32 ECG + MAX30102 PPG monitoring with Web Bluetooth dashboard
Current build compiles for ESP32 Dev Module
PPG waveform is raw IR; ECG is filtered before transmission
Arrhythmia classification is heuristic, not medical grade
Firmware
ESP32 samples ECG, MAX30102 red/IR, computes vitals, sends BLE notifications.
Dashboard
Mobile web app connects over BLE, renders waveforms, and stores history locally.
Display
16x2 LCD shows PPG schedule, BPM, SpO2, and sensor status.
HeartRateAnalyze | Slide 1/11

## Slide 2
Problem & Goal
A compact vital-sign monitor for demo and learning
Capture ECG and optical pulse data from low-cost sensors.
Show live heart rate, SpO2 estimate, HRV, rhythm status, ECG, and PPG.
Use browser-based UI so no native mobile app is required.
Persist short recording history in localStorage for review.
HeartRateAnalyze | Slide 2/11

## Slide 3
System Architecture
End-to-end data flow
MAX30102 red/IR feeds BPM, SpO2 estimate, and raw PPG waveform.
ECG pin 34 feeds filtered ECG waveform.
BLE service uses fixed UUIDs matched by the browser.
ECG Sensor
ESP32 Firmware
BLE Notify
Web App
Charts + History
HeartRateAnalyze | Slide 3/11

## Slide 4
Hardware Components
Main electronics and pin mapping
Libraries: BLEDevice/BLEServer, SparkFun MAX30105, heartRate, LiquidCrystal_I2C
ESP32
BLE peripheral named VitalMonitor
Analog ECG input on GPIO 34
I2C on GPIO 21/22
MAX30102
Raw IR PPG waveform
Red + IR values for SpO2
Beat detection from IR channel
16x2 LCD
I2C address 0x27
PPG schedule/recording status
BPM and SpO2 display
HeartRateAnalyze | Slide 4/11

## Slide 5
Firmware Workflow
Loop behavior on the ESP32
ECG is sampled at 250 Hz from GPIO 34.
ECG passes through high-pass baseline removal and low-pass smoothing.
MAX30102 is read once per PPG tick to avoid FIFO/read conflicts.
Raw IR is sent as the PPG waveform; red + IR are used internally for SpO2.
Vitals JSON is sent once per second over BLE.
HeartRateAnalyze | Slide 5/11

## Slide 6
Signal Processing
What is processed and what stays raw
ECG DSP
High-pass style baseline removal
Low-pass smoothing
Normalized into 0..4095 for graphing
Raw PPG
MAX30102 IR is sent directly
No DSP before BLE
Used for live PPG waveform
SpO2
Rolling red/IR window
Ratio-of-ratios estimate
Invalid signal sends null
HeartRateAnalyze | Slide 6/11

## Slide 7
Arrhythmia Classification
Heuristic rhythm labeling from BPM and RR variability
NO_SIGNAL: insufficient beat/RR data.
NORMAL: BPM and variability within simple thresholds.
BRADYCARDIA: BPM below 45.
TACHYCARDIA: BPM above 120.
IRREGULAR: high RMSSD variability.
LOW_HRV: RMSSD below threshold.
HeartRateAnalyze | Slide 7/11

## Slide 8
BLE Contract
Device and characteristic interface
Device name: VitalMonitor
Service: 12345678-1234-1234-1234-123456789abc
Vitals characteristic sends JSON: bpm, spo2, hrv, rhythm, alert.
ECG and PPG characteristics send comma-separated numeric chunks.
Alert characteristic sends ARRHYTHMIA_ALERT when abnormal rhythm is detected.
HeartRateAnalyze | Slide 8/11

## Slide 9
Web Dashboard Workflow
Browser connects, subscribes, draws, and stores
User taps Connect to Device in Chrome with Web Bluetooth support.
App subscribes to vitals, ECG, PPG, and alert characteristics.
Canvas renderers draw live ECG and PPG waveforms.
Each vitals packet creates a history record with recent ECG/PPG samples.
History is stored in localStorage with filters for alerts, today, and week.
HeartRateAnalyze | Slide 9/11

## Slide 10
Current Status & Testing
What is working now
Firmware compiles for esp32:esp32:esp32.
LCD shows MAX30102 status, PPG schedule/recording, BPM, and SpO2.
PPG sensor failure no longer freezes silently; LCD shows error.
Compile size: about 86% program storage, 13% dynamic memory.
Next hardware checks: upload, serial monitor, BLE connect, live waveform validation.
HeartRateAnalyze | Slide 10/11

## Slide 11
Limitations & Next Steps
Known constraints and future improvements
SpO2 and rhythm classification are demo heuristics, not medical-grade algorithms.
MAX30102 readings depend heavily on finger placement, light leakage, and wiring.
Improve ECG classification with real QRS detection and tested RR interval windows.
Add calibration/debug screens for raw red/IR values and signal quality.
Add export/share option for recorded sessions from the web dashboard.
HeartRateAnalyze | Slide 11/11